Dear Yoke,

I have been looking at the current clamp mode of the Pico and have encountered some oddities in scaling and calibration. I am also getting a hardware buffer overflow error that might be worth investigating.

Gain in ICLAMP modes

In the ICLAMP modes, the highest voltage gain setting (X80) does not seem to be operating. When it is selected the measured resistance is half what it should be. It appears that the gain at X80 is the next highest in the list (X40).

| Pico 2: Effect of varying gain in ICLAMP (255pA) mode on holding potential and resistance |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Head (100 M res.) |  | Imodel |  |
| Gain | V hold | Resistance | V hold | Resistance |
| 1 | 4.7 mV | 102 M | 0.6 | 9 M |
| 10 | 4.6 | 100.5 M | 0.4 | 8.9 M |
| 20 | 4.7 | 100 M | 0.4 | 8.8 M |
| 40 | 4.5 | 98.5 M | 0.3 | 8.8 M |
| 80 | 2.3 | 49 M | 0.2 | 4.4 M |

Current calibration in ICLAMP modes

The current stimulus calibration seems to vary between ICLAMP modes. This is most noticeable when using a 100 M resistor attached to the head stage, compared to the IMODEL sources.

|  | Pico 2 |  | Pico 1 |  |
| --- | --- | --- | --- | --- |
|  | Head (100 M) | Model | Head (100M) | Model |
| VClamp (255mV) | 100.3 M | 109.3 M | 102.8 M | 112.6 M |
| VClamp (2.048V) | 98.3 M | 106.9 M | 102.1 M | 111.9 M |
| VClamp (4.096V) | - | - | No pulse | No pulse |
| IClamp (255 pA) 20X | 99.5 M | 8.8 M | 104.5 M | 1.04 M |
| IClamp (2.55 nA) 20X | 131.6 M | 9.9 M | 121.1 M | 1.13 M |
| IClamp (25.5 nA) 20X | Oscillates | 10.9 M | 122.9 M | 1.03 M |

Note how the 100 M resistance is measured by the Pico 2 as 131.6 M in IClamp (2.55 nA) but 99.5 M in IClamp (255 pA). There is some variance in the IMODEL measurements but not as much.

Voltage offset in ICLAMP mode

An unexplained voltage offset appears on the headstage voltage channel when in ICLAMP mode. With a 100 M resistor connected to the headstage, an offset of 4.3 mV appears with the Pico 2 (and 21.5 mV with the Pico 1) in ICLAMP mode with the holding current set to zero. This is not consistent with the holding current measured in VCLAMP mode with the holding voltage set to zero. This offset in ICLAMP mode is not observed when using the IMODEL resistor as the input source.

|  | Pico 2 |  | Pico 1 |  |
| --- | --- | --- | --- | --- |
| Mode | ICLAMP(255) | VCLAMP(255) | ICLAMP(255) | VCLAMP(255) |
|  | Holding Pot. | Holding Current | Holding Pot. | Holding Current |
| Open | 0.0 mV | 0.0 pA | -0.5 mV | -10 pA |
| Head (with 100 M resistor) | 4.3 mV | 1 pA | 21.5 mV | 8 pA |
| IModel | 0.4 mV | 2 pA | -0.3 mV | 7 pA |

Is there an offset in the headstage that is not being calibrated for correctly? Or could it be due to the higher value of the resistor attached to the headstage?

Hardware buffer overflow errors

I have been getting intermittent “hardware buffer overflow” errors (error 769) when I call tecella_auto_scale and also tecella_acquire_read_i (resulting in the loss of data for the stimulus sweep after the error occurs). The errors do not occur consistently. They occurs most regularly on a 4 year old Dell D630 laptop I use for development and much less often on the faster and newer desktop that I use. However, some days it does not occur at all on the laptop. When developing the support for streaming stimuli, I discovered that the error occurred frequently during tecella_acquire_read_i until I extended the stimulus waveform beyond the end of the sweep by at least 0.25 second (i.e. for a 100 ms sweep, I create a 350 ms stimulus). At present (with WinWCP V4.4.6) the error continues to occur intermittently with the older segmented stimulus mode but not with the streaming stimulus.

I have been working round the errors for some time, but I think it might be worth trying to resolve them. Can you tell me what causes a hardware buffer overflow and is there anything that can be done to avoid the errors completely?