Installing WCP for Windows with the CED 1401, John Dempster 3/4/97

INSTALLATION

1) Install the CED 1401 driver and 1401 commands, using the INSTALL.EXE program on the CED 1401 Installation Disk supplied by Cambridge Electronic Design. (This installs the device driver CED1401.SYS and creates the directory \1401).

2) Install CED 1401 Windows support software, using the SETUP.EXE program on the disk supplied by Cambridge Electronic Design. (This adds a CED 1401 icon to the Windows Control Panel, and other files which support the use of the CED 1401 from a Windows program).

3) Test the CED 1401 using TRY1401.EXE.

4) Install the WCP for Windows software, using the SETUP.EXE program on the WCP for Windows installation disk. (This installs WinWCP in the directory \winwcp and creates a Strathclyde Software program group.

CONNECTIONS

A/D Inputs

WCP supports up to 6 A/D input channels (4 on Micro1401); ADC 0, ADC 1, ADC 2, ADC 3. (all 1401s), ADC 5, ADC 6 (1401, 1401+).

D/A Outputs

Voltage clamp command voltage signals are output on DAC 0.

A synchronisation pulse which triggers the start of a recording sweep is supplied on DAC 1. The trigger signal is a 5V, 1ms duration, active low pulse.

Digital Outputs

Digital On/Off pulse patterns can be be set to digital output bits 8-15 (bits 8,9 appear on the Digital Outputs 0 and 1 on the Micro1401 front panel), synchronously with the command voltage waveform.

Triggering

WinWCP, in Pulse Program mode, generates command voltage waveforms on DAC 0, digital pulses on digital outputs 8-15, and a pulse on DAC 1 which triggers the recording sweep.

DAC 1 should be connected to Event Input 4 (labelled Trigger Input on the Micro1401) and also to Event Input 2 (Pin 3 on the rear 15 pin D-socket on the Micro1401).

(Note the connection to Event Input 2 is not required if digital output patterns are not being used.)

UPDATES

The latest versions of the software can be obtained from my anonymous FTP site

ppserver.dpp.strath.ac.uk

The file WINWCP??.EXE (where ?? is the version number) in the directory /pub/ses contains the latest version of the software.